## Measurements
Identification |  |  | Router TMCM (HEX) | TMCM Router [°C] | RMCM Router [Ω] | Router Tbus (HEX) | Thermal-Camera | RBUS Router [Ω] | IAPILOT [mA] | TPt1000 [°C] | TOffset [°C] | RBus Pt1000 [Ω] | ΔRBUS [Ω] | T trhreshold [°C] | Threshold for Router (HEX) | THR +5°C | THR -5°C | TBUS Router [°C] | Comments | IBus [A]
Sector | Side | Half-Stave
0 | C | CH0 | 26E | 20.7149 | 5578.2833 | 267 | 28 | 5539 | 0.24573 | 29.6 | -1.6 | 5569.8 | 30.8 | 35 | 291 | 279 | 256 | 28 |  | 4.2 | 657 | 1419.003
|  | CH1 | 2A0 | 19.364 | 5894.0237 | 25C | 26.5 | 5510.125 | 0.24426 | 27.7 | -1.2 | 5533.225 | 23.1 | 35 | 28B | 26E | 24B | 26.5 |  | 4.58 | 651 | 1410.729
|  | CH2 | 298 | 19.4111 | 5882.6712 | 253 | 25.5 | 5490.875 | 0.24286 | 26.5 | -1 | 5510.125 | 19.25 | 35 | 285 | 265 | 242 | 25.5 |  | 3.63 | 645 | 1402.455
|  | CH3 | 285 | 20.3484 | 5661.9505 | 268 | 26 | 5500.5 | 0.2477 | 27.8 | -1.8 | 5535.15 | 34.65 | 35 | 29A | 27A | 257 | 26 |  | 4.8 | 666 | 1431.414
|  | CH4 | 27F | 20.2975 | 5673.6857 | 267 | 28 | 5539 | 0.24573 | 29.3 | -1.3 | 5564.025 | 25.025 | 35 | 291 | 279 | 256 | 28 |  | 4.3 | 657 | 1419.003
|  | CH5 | 28D | 20.0383 | 5733.8844 | 265 | 26.5 | 5510.125 | 0.24651 | 27.8 | -1.3 | 5535.15 | 25.025 | 35 | 295 | 277 | 254 | 26.5 |  | 4.69 | 661 | 1424.519
| A | CH0 | 2BE | 18.9286 | 6000.2098 | 268 | 27 | 5519.75 | 0.24683 | 29.2 | -2.2 | 5562.1 | 42.35 | 35 | 296 | 27A | 257 | 27 |  | 4.81 | 662 | 1425.898
|  | CH1 | 2C8 | 18.4706 | 6114.3289 | 262 | 28 | 5539 | 0.24448 | 27.9 | 0.1 | 5537.075 | -1.925 | 35 | 28C | 274 | 251 | 28 |  | 4.01 | 652 | 1412.108
|  | CH2 | 2AD | 19.2186 | 5929.24 | 264 | 27 | 5519.75 | 0.24584 | 29.3 | -2.3 | 5564.025 | 44.275 | 35 | 292 | 276 | 253 | 27 |  | 5.32 | 658 | 1420.382
|  | CH3 | [19 empty] | 513
|  | CH4 | 2A0 | 19.3556 | 5896.069 | 26B | 31 | 5596.75 | 0.24418 | 29.55 | 1.45 | 5568.8375 | -27.9125 | 35 | 28B | 27D | 25B | 31 |  | 4.4 | 651 | 1410.729
|  | CH5 | 2AC | 19.2311 | 5926.2215 | 267 | 28 | 5539 | 0.24573 | 28.55 | -0.55 | 5549.5875 | 10.5875 | 35 | 291 | 279 | 256 | 28 |  | 4.86 | 657 | 1419.003
1 | C | CH0 | 269 | 20.8392 | 5550.2238 | 267 | 28 | 5539 | 0.24573 | 32.7 | -4.7 | 5629.475 | 90.475 | 35 | 291 | 279 | 256 | 28 |  | 5.19 | 657 | 1419.003
|  | CH1 | 290 | 19.684 | 5817.3984 | 258 | 26 | 5500.5 | 0.24369 | 31.8 | -5.8 | 5612.15 | 111.65 | 35 | 289 | 26A | 247 | 26 |  | 5.22 | 649 | 1407.971
|  | CH2 | 29E | 19.1169 | 5954.0121 | 252 | 27 | 5519.75 | 0.24134 | 28.9 | -1.9 | 5556.325 | 36.575 | 35 | 27F | 264 | 242 | 27 |  | 5.23 | 639 | 1394.181
|  | CH3 | 2AB | 19.3861 | 5888.712 | 262 | 25 | 5481.25 | 0.24706 | 28.8 | -3.8 | 5554.4 | 73.15 | 35 | 297 | 274 | 251 | 25 |  | 5.34 | 663 | 1427.277
|  | CH4 | 291 | 19.8147 | 5786.4252 | 265 | 28 | 5539 | 0.24523 | 31.6 | -3.6 | 5608.3 | 69.3 | 35 | 28F | 277 | 254 | 28 |  | 5.38 | 655 | 1416.245
|  | CH5 | 292 | 19.766 | 5797.9347 | 264 | 28 | 5539 | 0.24498 | 30.7 | -2.7 | 5590.975 | 51.975 | 35 | 28E | 276 | 253 | 28 |  | 5.18 | 654 | 1414.866
| A | CH0 | 2CD | 18.4576 | 6117.6125 | 266 | 28 | 5539 | 0.24548 | 32.6 | -4.6 | 5627.55 | 88.55 | 35 | 290 | 278 | 255 | 28 |  | 4.88 | 656 | 1417.624
|  | CH1 | 2BE | 18.6608 | 6066.622 | 264 | 29 | 5558.25 | 0.24413 | 31.7 | -2.7 | 5610.225 | 51.975 | 35 | 28B | 276 | 254 | 29 |  | 5.03 | 651 | 1410.729
|  | CH2 | 29B | 19.6599 | 5823.1245 | 25E | 25 | 5481.25 | 0.24605 | 29.6 | -4.6 | 5569.8 | 88.55 | 35 | 293 | 270 | 24D | 25 |  | 5.39 | 659 | 1421.761
|  | CH3 | 277 | 20.328 | 5666.6407 | 25D | 27 | 5519.75 | 0.24409 | 30.8 | -3.8 | 5592.9 | 73.15 | 35 | 28B | 26F | 24C | 27 |  | 5.34 | 651 | 1410.729
|  | CH4 | 29F | 19.355 | 5896.2159 | 259 | 26 | 5500.5 | 0.24394 | 31.4 | -5.4 | 5604.45 | 103.95 | 35 | 28A | 26B | 248 | 26 |  | 5.42 | 650 | 1409.35
|  | CH5 | 2AC | 19.1816 | 5938.2543 | 265 | 28 | 5539 | 0.24523 | 31.1 | -3.1 | 5598.675 | 59.675 | 35 | 28F | 277 | 254 | 28 |  | 5.29 | 655 | 1416.245
2 | C | CH0 | 240 | 21.8178 | 5335.0687 | 268 | 29.1 | 5560.175 | 0.24504 | 28.9 | 0.2 | 5556.325 | -3.85 | 35 | 28F | 27A | 257 | 29.1 |  | 5.44 | 655 | 1416.245
|  | CH1 | 263 | 20.8861 | 5539.6977 | 268 | 29.5 | 5567.875 | 0.2447 | 27.6 | 1.9 | 5531.3 | -36.575 | 35 | 28D | 27A | 258 | 29.5 |  | 5.31 | 653 | 1413.487
|  | CH2 | 277 | 20.8121 | 5556.325 | 277 | 28.9 | 5556.325 | 0.24893 | 29.6 | -0.7 | 5569.8 | 13.475 | 35 | 29F | 289 | 266 | 28.9 |  | 5.34 | 671 | 1438.309
|  | CH3 | 24E | 21.8204 | 5334.5175 | 276 | 28.9 | 5556.325 | 0.24868 | 29.21 | -0.31 | 5562.2925 | 5.9675 | 35 | 29E | 288 | 265 | 28.9 |  | 5.34 | 670 | 1436.93
|  | CH4 | 268 | 20.9898 | 5516.4747 | 26B | 27.7 | 5533.225 | 0.24698 | 27.16 | 0.54 | 5522.83 | -10.395 | 35 | 297 | 27D | 25A | 27.7 |  | 5.53 | 663 | 1427.277
|  | CH5 | 25C | 21.0656 | 5499.5568 | 263 | 28 | 5539 | 0.24473 | 27.53 | 0.47 | 5529.9525 | -9.0475 | 35 | 28D | 275 | 252 | 28 |  | 5.39 | 653 | 1413.487
| A | CH0 | 29B | 19.8911 | 5768.4107 | 27A | 30.4 | 5585.2 | 0.24839 | 29.3 | 1.1 | 5564.025 | -21.175 | 35 | 29D | 28C | 269 | 30.4 |  | 5.88 | 669 | 1435.551
|  | CH1 | 28C | 19.9222 | 5761.0844 | 26A | 29.6 | 5569.8 | 0.24511 | 27.6 | 2 | 5531.3 | -38.5 | 35 | 28F | 27C | 259 | 29.6 |  | 5.25 | 655 | 1416.245
|  | CH2 | 23F | 22.2132 | 5250.9045 | 274 | 28.3 | 5544.775 | 0.2487 | 28.1 | 0.2 | 5540.925 | -3.85 | 35 | 29E | 286 | 263 | 28.3 |  | 5.2 | 670 | 1436.93
|  | CH3 | 228 | 22.454 | 5200.417 | 26B | 30 | 5577.5 | 0.24502 | 29.1 | 0.9 | 5560.175 | -17.325 | 35 | 28E | 27D | 25B | 30 |  | 5.2 | 654 | 1414.866
|  | CH4 | 231 | 22.0903 | 5276.9083 | 25F | 27.9 | 5537.075 | 0.24382 | 27.84 | 0.06 | 5535.92 | -1.155 | 35 | 289 | 271 | 24E | 27.9 |  | 5.34 | 649 | 1407.971
|  | CH5 | 247 | 21.8276 | 5332.9736 | 265 | 26 | 5500.5 | 0.24695 | 26.51 | -0.51 | 5510.3175 | 9.8175 | 35 | 297 | 277 | 254 | 26 |  | 5.34 | 663 | 1427.277
3 | C | CH0 | 29F | 19.499 | 5861.5675 | 268 | 28.7 | 5552.475 | 0.24538 | 32.7 | -4 | 5629.475 | 77 | 35 | 290 | 27A | 257 | 28.7 |  | 5.38 | 656 | 1417.624
|  | CH1 | 287 | 19.9364 | 5757.7574 | 272 | 33.2 | 5639.1 | 0.24406 | 35.6 | -2.4 | 5685.3 | 46.2 | 35 | 28A | 284 | 262 | 33.2 |  | 4.83 | 650 | 1409.35
|  | CH2 | 26F | 20.7593 | 5568.2327 | 271 | 30.1 | 5579.425 | 0.24642 | 31.85 | -1.75 | 5613.1125 | 33.6875 | 35 | 294 | 283 | 260 | 30.1 |  | 5.18 | 660 | 1423.14
|  | CH3 | 255 | 21.5833 | 5385.7355 | 273 | 28.7 | 5552.475 | 0.24811 | 33.6 | -4.9 | 5646.8 | 94.325 | 35 | 29B | 285 | 262 | 28.7 |  | 5.4 | 667 | 1432.793
|  | CH4 | 1AD | 26.1162 | 4498.31 | 267 | 28.2 | 5542.85 | 0.24556 | 32.09 | -3.89 | 5617.7325 | 74.8825 | 35 | 291 | 279 | 256 | 28.2 |  | 5.06 | 657 | 1419.003
|  | CH5 | 267 | 20.7056 | 5580.3807 | 25F | 27.8 | 5535.15 | 0.24391 | 30.87 | -3.07 | 5594.2475 | 59.0975 | 35 | 28A | 271 | 24E | 27.8 |  | 5.16 | 650 | 1409.35
| A | CH0 | 2BA | 18.9037 | 6006.3494 | 264 | 27.2 | 5523.6 | 0.24566 | 31.3 | -4.1 | 5602.525 | 78.925 | 35 | 291 | 276 | 253 | 27.2 |  | 5 | 657 | 1419.003
|  | CH1 | 266 | 21.5251 | 5398.3761 | 299 | 35.2 | 5677.6 | 0.25187 | 41.2 | 0 | 5428 | -249.6 | 35 | 2AB | 2AC | 288 | 35.2 |  | 5.41 | 683 | 1454.857
|  | CH2 | 290 | 19.9748 | 5748.7478 | 270 | 29.6 | 5569.8 | 0.2466 | 32.6 | -3 | 5627.55 | 57.75 | 35 | 295 | 282 | 25F | 29.6 |  |  | 661 | 1424.519
|  | CH3 | 25D | 21.2407 | 5460.7504 | 26B | 28 | 5539 | 0.24672 | 31.6 | -3.6 | 5608.3 | 69.3 | 35 | 296 | 27D | 25A | 28 |  | 5.1 | 662 | 1425.898
|  | CH4 | 2B0 | 19.0666 | 5966.3319 | 261 | 27.1 | 5521.675 | 0.245 | 30.24 | -3.14 | 5582.12 | 60.445 | 35 | 28E | 273 | 250 | 27.1 |  | 5.3 | 654 | 1414.866
|  | CH5 | 28A | 19.9054 | 5765.0379 | 265 | 28.9 | 5556.325 | 0.24447 | 32.47 | -3.57 | 5625.0475 | 68.7225 | 35 | 28C | 277 | 254 | 28.9 |  | 5.1 | 652 | 1412.108
4 | C | CH0 | 276 | 20.5479 | 5616.2247 | 26C | 29.1 | 5560.175 | 0.24603 | 29.53 | -0.43 | 5568.4525 | 8.2775 | 35 | 293 | 27E | 25B | 29.1 |  | 5.16 | 659 | 1421.761
|  | CH1 | 279 | 19.3951 | 5886.5341 | 255 | 35.1 | 5675.675 | 0.23544 | 35.32 | -0.22 | 5679.91 | 4.235 | 35 | 266 | 267 | 245 | 35.1 | check | 5.17 | 614 | 1359.706
|  | CH2 | 258 | 21.2233 | 5464.6015 | 269 | 29.1 | 5560.175 | 0.24529 | 29.49 | -0.39 | 5567.6825 | 7.5075 | 35 | 290 | 27B | 258 | 29.1 |  | 5.36 | 656 | 1417.624
|  | CH3 | 210 | 23.1868 | 5050.2032 | 26F | 30.3 | 5583.275 | 0.24575 | 30.26 | 0.04 | 5582.505 | -0.77 | 35 | 292 | 281 | 25E | 30.3 |  | 5.14 | 658 | 1420.382
|  | CH4 | 26C | 14.4467 | 7233.5626 | 18B | 30.8 | 5592.9 | 0.18912 | 31.07 | -0.27 | 5598.0975 | 5.1975 | 35 | 1A4 | 199 | 17E | 30.8 |  | 4.55 | 420 | 1092.18
|  | CH5 | 266 | 20.7354 | 5573.65 | 266 | 29.8 | 5573.65 | 0.24395 | 30.02 | -0.22 | 5577.885 | 4.235 | 35 | 28A | 278 | 256 | 29.8 |  | 4.13 | 650 | 1409.35
| A | CH0 | 2DD | 18.2526 | 6169.5219 | 26A | 27.4 | 5527.45 | 0.24699 | 27.06 | 0.34 | 5520.905 | -6.545 | 35 | 297 | 27C | 259 | 27.4 |  | 5.35 | 663 | 1427.277
|  | CH1 | hot !!! |  | #NUM! | 253 | 21 | 5404.25 | 0.24675 |  |  | 5424 | 19.75 | 35 | 296 | 265 | 242 | 21 |  | 5.38 | 662 | 1425.898
|  | CH2 | 27C | 20.162 | 5705.0572 | 25F | 28.1 | 5540.925 | 0.24365 | 27.75 | 0.35 | 5534.1875 | -6.7375 | 35 | 289 | 271 | 24F | 28.1 |  | 5.23 | 649 | 1407.971
|  | CH3 | 209 | 23.1548 | 5056.6624 | 266 | 30.3 | 5583.275 | 0.24353 | 30.48 | -0.18 | 5586.74 | 3.465 | 35 | 288 | 278 | 256 | 30.3 |  | 5.24 | 648 | 1406.592
|  | CH4 | 224 | 23.2505 | 5037.3994 | 267 | 21 | 5404.25 | 0.25185 | 31.08 | -10.08 | 5554 | 149.75 | 35 | 2AB | 279 | 256 | 21 |  | 5.26 | 683 | 1454.857
|  | CH5 | 258 | 20.9298 | 5529.8902 | 25E | 29.3 | 5564.025 | 0.24239 | 29.34 | -0.04 | 5564.795 | 0.77 | 35 | 283 | 270 | 24E | 29.3 |  | 5.17 | 643 | 1399.697
5 | C | CH0 | 266 | 20.5746 | 5610.1397 | 265 | 31.4 | 5604.45 | 0.24237 | 29.72 | 1.68 | 5572.11 | -32.34 | 35 | 283 | 277 | 255 | 31.4 |  | 4.81 | 643 | 1399.697
|  | CH1 | 26D | 20.6098 | 5602.1247 | 26A | 30.4 | 5585.2 | 0.24444 | 28.65 | 1.75 | 5551.5125 | -33.6875 | 35 | 28C | 27C | 25A | 30.4 |  | 5.2 | 652 | 1412.108
|  | CH2 | 260 | 20.8216 | 5554.185 | 25E | 28.2 | 5542.85 | 0.24332 | 28.72 | -0.52 | 5552.86 | 10.01 | 35 | 287 | 270 | 24E | 28.2 |  | 5.31 | 647 | 1405.213
|  | CH3 | 23F | 22.0117 | 5293.6037 | 26E | 28.9 | 5556.325 | 0.2467 | 29 | -0.1 | 5558.25 | 1.925 | 35 | 296 | 280 | 25D | 28.9 |  | 4.92 | 662 | 1425.898
|  | CH4 | 23D | 21.9947 | 5297.2261 | 26F | 30 | 5577.5 | 0.24601 | 29.7 | 0.3 | 5571.725 | -5.775 | 35 | 293 | 281 | 25E | 30 |  | 5 | 659 | 1421.761
|  | CH5 | 24E | 21.2409 | 5460.7145 | 268 | 31.6 | 5608.3 | 0.24294 | 31.56 | 0.04 | 5607.53 | -0.77 | 35 | 286 | 27A | 258 | 31.6 |  | 5.18 | 646 | 1403.834
| A | CH0 | 2B0 | 18.9385 | 5997.7845 | 263 | 29.2 | 5562.1 | 0.24372 | 27.87 | 1.33 | 5536.4975 | -25.6025 | 35 | 289 | 275 | 253 | 29.2 |  | 4.99 | 649 | 1407.971
|  | CH1 | 2CA | 18.3175 | 6153.0435 | 260 | 28.7 | 5552.475 | 0.24339 | 28.28 | 0.42 | 5544.39 | -8.085 | 35 | 288 | 272 | 250 | 28.7 |  | 4.96 | 648 | 1406.592
|  | CH2 | 26F | 20.7625 | 5567.5159 | 26D | 28.9 | 5556.325 | 0.24645 | 29.39 | -0.49 | 5565.7575 | 9.4325 | 35 | 294 | 27F | 25C | 28.9 |  | 5.38 | 660 | 1423.14
|  | CH3 | 26D | 20.9019 | 5536.1494 | 272 | 29.3 | 5564.025 | 0.24735 | 30.24 | -0.94 | 5582.12 | 18.095 | 35 | 298 | 284 | 261 | 29.3 |  | 5.24 | 664 | 1428.656
|  | CH4 | 272 | 20.2033 | 5695.487 | 25E | 30.2 | 5581.35 | 0.24164 | 30.22 | -0.02 | 5581.735 | 0.385 | 35 | 280 | 270 | 24E | 30.2 |  | 4.99 | 640 | 1395.56
|  | CH5 | 280 | 20.0458 | 5732.1344 | 263 | 29.5 | 5567.875 | 0.24346 | 30.08 | -0.58 | 5579.04 | 11.165 | 35 | 288 | 275 | 253 | 29.5 |  | 5.14 | 648 | 1406.592
6 | C | CH0 | 22C | 22.3576 | 5220.5671 | 267 | 28.7 | 5552.475 | 0.24513 | 31 | -2.3 | 5596.75 | 44.275 | 35 | 28F | 279 | 256 | 28.7 |  | 5.69 | 655 | 1416.245
|  | CH1 | 23D | 21.72 | 5356.1337 | 261 | 29.1 | 5560.175 | 0.2433 | 31.7 | -2.6 | 5610.225 | 50.05 | 35 | 287 | 273 | 251 | 29.1 |  | 5.19 | 647 | 1405.213
|  | CH2 | 25C | 21.1531 | 5480.1258 | 261 | 26.4 | 5508.2 | 0.2456 | 30.5 | -4.1 | 5587.125 | 78.925 | 35 | 291 | 273 | 250 | 26.4 |  | 4.77 | 657 | 1419.003
|  | CH3 | 25C | 21.2464 | 5459.4886 | 26E | 29.1 | 5560.175 | 0.24653 | 31.9 | -2.8 | 5614.075 | 53.9 | 35 | 295 | 280 | 25D | 29.1 |  | 5.16 | 661 | 1424.519
|  | CH4 | 21B | 22.9847 | 5091.119 | 271 | 29.7 | 5571.725 | 0.24676 | 31.7 | -2 | 5610.225 | 38.5 | 35 | 296 | 283 | 260 | 29.7 |  | 5.27 | 662 | 1425.898
|  | CH5 | 251 | 21.5537 | 5392.1697 | 26D | 28.5 | 5548.625 | 0.24679 | 32.5 | -4 | 5625.625 | 77 | 35 | 296 | 27F | 25C | 28.5 |  | 5.23 | 662 | 1425.898
| A | CH0 | 283 | 19.9123 | 5763.4289 | 253 | 25.5 | 5490.875 | 0.24286 | 29.4 | -3.9 | 5565.95 | 75.075 | 35 | 285 | 265 | 242 | 25.5 |  | 4.84 | 645 | 1402.455
|  | CH1 | 293 | 19.6948 | 5814.8163 | 25E | 26.8 | 5515.9 | 0.24451 | 30.15 | -3.35 | 5580.3875 | 64.4875 | 35 | 28C | 270 | 24D | 26.8 |  | 5.46 | 652 | 1412.108
|  | CH2 | 263 | 20.9282 | 5530.2559 | 25C | 25.5 | 5490.875 | 0.24512 | 29.5 | -4 | 5567.875 | 77 | 35 | 28F | 26E | 24B | 25.5 |  | 5.13 | 655 | 1416.245
|  | CH3 | 27D | 20.5666 | 5611.9501 | 26A | 26.3 | 5506.275 | 0.24794 | 31.2 | -4.9 | 5600.6 | 94.325 | 35 | 29B | 27C | 259 | 26.3 |  | 5.07 | 667 | 1432.793
|  | CH4 | 244 | 21.4902 | 5405.9869 | 256 | 26.4 | 5508.2 | 0.24285 | 29.9 | -3.5 | 5575.575 | 67.375 | 35 | 285 | 268 | 245 | 26.4 |  | 4.9 | 645 | 1402.455
|  | CH5 | 267 | 20.5876 | 5607.1717 | 25B | 28 | 5539 | 0.24274 | 30 | -2 | 5577.5 | 38.5 | 35 | 285 | 26D | 24B | 28 |  | 4.69 | 645 | 1402.455
7 | C | CH0 | 27D | 20.2891 | 5675.6237 | 263 | 27.5 | 5529.375 | 0.24516 | 32 | -4.5 | 5616 | 86.625 | 35 | 28F | 275 | 252 | 27.5 |  | 5.08 | 655 | 1416.245
|  | CH1 | 2A3 | 19.2662 | 5917.6834 | 260 | 28 | 5539 | 0.24398 | 32.7 | -4.7 | 5629.475 | 90.475 | 35 | 28A | 272 | 24F | 28 |  | 5.31 | 650 | 1409.35
|  | CH2 | 28F | 19.9439 | 5755.996 | 26A | 28.5 | 5548.625 | 0.24605 | 34 | -5.5 | 5654.5 | 105.875 | 35 | 293 | 27C | 259 | 28.5 |  | 5.1 | 659 | 1421.761
|  | CH3 | 28C | 19.852 | 5777.6304 | 260 | 27.5 | 5529.375 | 0.24441 | 33.2 | -5.7 | 5639.1 | 109.725 | 35 | 28C | 272 | 24F | 27.5 |  | 5.11 | 652 | 1412.108
|  | CH4 | 278 | 20.0937 | 5720.9552 | 253 | 26.5 | 5510.125 | 0.24201 | 31.6 | -5.1 | 5608.3 | 98.175 | 35 | 282 | 265 | 243 | 26.5 |  | 3.9 | 642 | 1398.318
|  | CH5 | 25A | 20.9954 | 5515.2137 | 265 | 30 | 5577.5 | 0.24354 | 33.22 | -3.22 | 5639.485 | 61.985 | 35 | 288 | 277 | 255 | 30 |  | 5.05 | 648 | 1406.592
| A | CH0 | 2FA | 17.6396 | 6327.8912 | 264 | 25.5 | 5490.875 | 0.24713 | 30.3 | -4.8 | 5583.275 | 92.4 | 35 | 297 | 276 | 253 | 25.5 |  | 5 | 663 | 1427.277
|  | CH1 | 2EB | 17.7402 | 6301.596 | 267 | 29 | 5558.25 | 0.24488 | 32.1 | -3.1 | 5617.925 | 59.675 | 35 | 28E | 279 | 256 | 29 |  | 5.08 | 654 | 1414.866
|  | CH2 | 2B4 | 19.075 | 5964.278 | 26F | 30 | 5577.5 | 0.24601 | 33.15 | -3.15 | 5638.1375 | 60.6375 | 35 | 293 | 281 | 25E | 30 |  | 5.25 | 659 | 1421.761
|  | CH3 | 294 | 19.6821 | 5817.8456 | 268 | 29.6 | 5569.8 | 0.24462 | 31.8 | -2.2 | 5612.15 | 42.35 | 35 | 28D | 27A | 258 | 29.6 |  | 5.4 | 653 | 1413.487
|  | CH4 | 2AF | 18.9189 | 6002.6106 | 264 | 30 | 5577.5 | 0.24329 | 31.6 | -1.6 | 5608.3 | 30.8 | 35 | 287 | 276 | 254 | 30 |  | 5.21 | 647 | 1405.213
|  | CH5 | 297 | 19.4381 | 5876.1859 | 25F | 29 | 5558.25 | 0.24289 | 30.8 | -1.8 | 5592.9 | 34.65 | 35 | 286 | 271 | 24F | 29 |  | 4.87 | 646 | 1403.834
8 | C | CH0 | 290 | 19.8797 | 5771.0931 | 267 | 28.1 | 5540.925 | 0.24564 | 32.8 | -4.7 | 5631.4 | 90.475 | 35 | 291 | 279 | 256 | 28.1 |  | 5.29 | 657 | 1419.003
|  | CH1 | 262 | 20.9895 | 5516.5296 | 266 | 28 | 5539 | 0.24548 | 34.7 | -6.7 | 5667.975 | 128.975 | 35 | 290 | 278 | 255 | 28 |  | 5.02 | 656 | 1417.624
|  | CH2 | 27B | 20.732 | 5574.3976 | 266 | 23.8 | 5458.15 | 0.24911 | 30.27 | -6.47 | 5582.6975 | 124.5475 | 35 | 2A0 | 278 | 255 | 23.8 |  | 4.87 | 672 | 1439.688
|  | CH3 | 2A1 | 19.1638 | 5942.5737 | 258 | 27.4 | 5527.45 | 0.2425 | 30.5 | -3.1 | 5587.125 | 59.675 | 35 | 284 | 26A | 248 | 27.4 |  | 4.44 | 644 | 1401.076
|  | CH4 | 263 | 21.0293 | 5507.6522 | 26B | 28.7 | 5552.475 | 0.24612 | 34.6 | -5.9 | 5666.05 | 113.575 | 35 | 293 | 27D | 25A | 28.7 |  | 5.01 | 659 | 1421.761
|  | CH5 | 254 | 21.1758 | 5475.0965 | 261 | 28.5 | 5548.625 | 0.24381 | 34.3 | -5.8 | 5660.275 | 111.65 | 35 | 289 | 273 | 251 | 28.5 |  | 4.81 | 649 | 1407.971
| A | CH0 | 31B | 17.0465 | 6485.6705 | 268 | 25.5 | 5490.875 | 0.24813 | 30.7 | -5.2 | 5590.975 | 100.1 | 35 | 29C | 27A | 257 | 25.5 |  | 4.79 | 668 | 1434.172
|  | CH1 | 2FF | 17.8323 | 6277.618 | 272 | 26 | 5500.5 | 0.25021 | 32.8 | -6.8 | 5631.4 | 130.9 | 35 | 2A4 | 284 | 261 | 26 |  | 4.44 | 676 | 1445.204
|  | CH2 | 2B6 | 18.6895 | 6059.4837 | 254 | 26.1 | 5502.425 | 0.2426 | 30.7 | -4.6 | 5590.975 | 88.55 | 35 | 284 | 266 | 243 | 26.1 |  | 4.63 | 644 | 1401.076
|  | CH3 | 2D0 | 18.3971 | 6132.882 | 25D | 25.3 | 5487.025 | 0.24554 | 30.3 | -5 | 5583.275 | 96.25 | 35 | 291 | 26F | 24C | 25.3 |  | 4.47 | 657 | 1419.003
|  | CH4 | 2AE | 18.922 | 6001.8326 | 258 | 26.7 | 5513.975 | 0.24309 | 30.2 | -3.5 | 5581.35 | 67.375 | 35 | 286 | 26A | 247 | 26.7 |  | 3.88 | 646 | 1403.834
|  | CH5 | 26B | 22.2027 | 5253.1196 | 2B2 | 32.7 | 5629.475 | 0.26015 | 32.7 | 0 | 5629.475 | 0 | 35 | 2CE | 2C5 | 2A1 | 32.7 |  | 5.26 | 718 | 1503.122
9 | C | CH0 | 2CE | 18.0428 | 6223.1975 | 248 | 23.8 | 5458.15 | 0.24154 | 25.3 | -1.5 | 5487.025 | 28.875 | 35 | 280 | 25A | 237 | 23.8 |  | 5.52 | 640 | 1395.56
|  | CH1 | 2CF | 18.4257 | 6125.6522 | 26C | 29.6 | 5569.8 | 0.24561 | 31.3 | -1.7 | 5602.525 | 32.725 | 35 | 291 | 27E | 25B | 29.6 |  | 5.42 | 657 | 1419.003
|  | CH2 | 2A6 | 19.3155 | 5905.7644 | 260 | 26.6 | 5512.05 | 0.24518 | 27.6 | -1 | 5531.3 | 19.25 | 35 | 28F | 272 | 24F | 26.6 |  | 5.48 | 655 | 1416.245
|  | CH3 | 25D | 21.1057 | 5490.6565 | 268 | 28.7 | 5552.475 | 0.24538 | 27.6 | 1.1 | 5531.3 | -21.175 | 35 | 290 | 27A | 257 | 28.7 |  | 5.3 | 656 | 1417.624
|  | CH4 | 2A6 | 19.0061 | 5981.1491 | 255 | 27 | 5519.75 | 0.24209 | 27.6 | -0.6 | 5531.3 | 11.55 | 35 | 282 | 267 | 245 | 27 |  | 5.38 | 642 | 1398.318
|  | CH5 | 2B8 | 18.4986 | 6107.2703 | 259 | 29.3 | 5564.025 | 0.24115 | 29.5 | -0.2 | 5567.875 | 3.85 | 35 | 27E | 26B | 249 | 29.3 |  | 5.48 | 638 | 1392.802
| A | CH0 | 2F3 | 17.8409 | 6275.3651 | 263 | 24.6 | 5473.55 | 0.24766 | 27.4 | -2.8 | 5527.45 | 53.9 | 35 | 29A | 275 | 252 | 24.6 |  | 5.49 | 666 | 1431.414
|  | CH1 | 316 | 17.1216 | 6465.4194 | 26B | 26.7 | 5513.975 | 0.24784 | 31.3 | -4.6 | 5602.525 | 88.55 | 35 | 29A | 27D | 25A | 26.7 |  | 5.35 | 666 | 1431.414
|  | CH2 | 2DF | 17.8323 | 6277.6056 | 258 | 26.6 | 5512.05 | 0.24318 | 27.9 | -1.3 | 5537.075 | 25.025 | 35 | 287 | 26A | 247 | 26.6 |  | 5.48 | 647 | 1405.213
|  | CH3 | 262 | 21.0347 | 5506.4427 | 263 | 26.6 | 5512.05 | 0.24593 | 28.2 | -1.6 | 5542.85 | 30.8 | 35 | 292 | 275 | 252 | 26.6 |  | 5.58 | 658 | 1420.382
|  | CH4 | 2E0 | 18.0996 | 6208.6261 | 263 | 26.4 | 5508.2 | 0.2461 | 28.2 | -1.8 | 5542.85 | 34.65 | 35 | 293 | 275 | 252 | 26.4 |  | 5.64 | 659 | 1421.761
|  | CH5 | 293 | 19.8494 | 5778.2369 | 269 | 28.2 | 5542.85 | 0.24605 | 30.8 | -2.6 | 5592.9 | 50.05 | 35 | 293 | 27B | 258 | 28.2 |  | 5.33 | 659 | 1421.761
10 | C | CH0 | 2E2 | 18.2024 | 6182.3297 | 26A | 26.7 | 5513.975 | 0.24759 | 30.4 | -3.7 | 5585.2 | 71.225 | 35 | 299 | 27C | 259 | 26.7 |  | 5.32 | 665 | 1430.035
|  | CH1 | 2A7 | 19.0396 | 5972.9283 | 25C | 28.4 | 5546.7 | 0.24265 | 31.9 | -3.5 | 5614.075 | 67.375 | 35 | 285 | 26E | 24C | 28.4 |  | 5.33 | 645 | 1402.455
|  | CH2 | 2C1 | 18.667 | 6065.0768 | 267 | 29 | 5558.25 | 0.24488 | 30.4 | -1.4 | 5585.2 | 26.95 | 35 | 28E | 279 | 256 | 29 |  | 5.48 | 654 | 1414.866
|  | CH3 | 2D5 | 18.4458 | 6120.5899 | 26A | 27.2 | 5523.6 | 0.24716 | 34.5 | -7.3 | 5664.125 | 140.525 | 35 | 297 | 27C | 259 | 27.2 |  | 5.5 | 663 | 1427.277
|  | CH4 | 2D8 | 18.107 | 6206.7329 | 262 | 28.1 | 5540.925 | 0.2444 | 31.1 | -3 | 5598.675 | 57.75 | 35 | 28C | 274 | 251 | 28.1 |  | 5.34 | 652 | 1412.108
|  | CH5 | 25B | 20.6948 | 5582.8334 | 252 | 27.6 | 5531.3 | 0.24083 | 32.4 | -4.8 | 5623.7 | 92.4 | 35 | 27D | 264 | 242 | 27.6 |  | 5.28 | 637 | 1391.423
| A | CH0 | 301 | 17.2577 | 6428.9629 | 261 | 27.4 | 5527.45 | 0.24474 | 29.5 | -2.1 | 5567.875 | 40.425 | 35 | 28D | 273 | 250 | 27.4 |  | 5.3 | 653 | 1413.487
|  | CH1 | 24F | 21.612 | 5379.4938 | 26F | 29 | 5558.25 | 0.24686 |  | 29 | 5000 | -558.25 | 35 | 296 | 281 | 25E | 29 |  | 5.35 | 662 | 1425.898
|  | CH2 | 2A0 | 18.7398 | 6046.9454 | 23E | 24.9 | 5479.325 | 0.23809 | 28.5 | -3.6 | 5548.625 | 69.3 | 35 | 271 | 24F | 22E | 24.9 |  | 5.57 | 625 | 1374.875
|  | CH3 | 2A9 | 19.146 | 5946.9107 | 25C | 26.6 | 5512.05 | 0.24418 | 30.3 | -3.7 | 5583.275 | 71.225 | 35 | 28B | 26E | 24B | 26.6 |  | 5.44 | 651 | 1410.729
|  | CH4 | 25B | 21.0903 | 5494.061 | 264 | 28.3 | 5544.775 | 0.24473 | 31.8 | -3.5 | 5612.15 | 67.375 | 35 | 28D | 276 | 253 | 28.3 |  | 5.34 | 653 | 1413.487
|  | CH5 | 288 | 23.333 | 5020.8592 | 2F4 | 28.7 | 5552.475 | 0.28015 | 31.8 | -3.1 | 5612.15 | 59.675 | 35 | 322 | 309 | 2E1 | 28.7 |  | 5.31 | 802 | 1618.958

## Legenda
Router TMCM (HEX) | Hexadecimal value given by the router and read out via LabView; The value is read when MCM and BUS is powered
Router TBUS (HEX) | Hexadecimal value given by the router and read out via LabView; The value is read when MCM and BUS is powered
Thermal- camera TC [°C] | Temperature measured with the Thermalcamera. The camera is pointed to the half-stave (= temperature of the BUS)
RMCM Router [Ω] | Calculated value of the MCM resistance in Ω: Formula: (1.379 * HEX2DEC(TMCM) + 513) / IAPILOT
RBUS Router [Ω] | Calculated value of the BUS resistance in Ω: Formula: (1.379 * HEX2DEC(TBUS) + 513) / IAPILOT
IAPILOT [mA] | Calculated value of the analogue pilot current in mA: Formula: (1.379 * HEX2DEC(TBUS) + 513) / ((TC * 19.25) + 5000)
TBUS Router [°C] | Calculated value of the BUS temperature in °C: Formula: (RBUSRouter - 5000) / 19.25
T MCM Router [°C] | Calculated value of the MCM temperature in °C: Formula: (3500 / (LN(RMCM Router / 4700) + 11.745)) - 273
TPt1000 [°C] | Temperature measured with the Pt1000 chain connected to the Cesar box in °C
TOffset [°C] | Difference between TPt1000 and TC in °C: Formula: TPt1000 - TC
RBus Pt1000 [Ω] | Calculated resistence of Pt1000 chain connected to the Cesa box in Ω: Formula: (TPt1000 * 3.85 + 1000) * 5
ΔR BUS [Ω] | Difference of the resistances between both Pt1000 chain (Router and Cesar box) in Ω: Formula: RBus Pt1000 - RBUS Router
THR +5°C | Calculated upper limit to set in the Router for the calibration of thePt1000 chain read by the router in hexadecimal: Formula: DEC2HEX((IAPILOT * (5000 + 19.5 * (TC + 5)) - 513) / 1.379)
THR -5°C | Calculated lower limit to set in the Router for the calibration of the Pt1000 chain read by the router in hexadecimal: Formula: DEC2HEX((IAPILOT * (5000 + 19.5 * (TC - 5)) - 513) / 1.379
Threshold for Router (HEX) | HExadecimal Threshold value to set on the Router for Interlock

## Statistical
I analog statistical
Class | Frequency | Comments | Mean | STD | Min | Max | %
0.189 | 0 |  | 0.24504 | 0.0063527 | 0.18912 | 0.28015 | 2.5925 |  | 0 | 0
0.19 | 1 | check on Sector 4 side C pos 4 | [6 empty] | 0 | 0
0.191 | 0 | [7 empty] | 0 | 0
0.192 | 0 | [7 empty] | 0 | 0
0.193 | 0 | [7 empty] | 0 | 0
0.194 | 0 | [7 empty] | 0 | 0
0.195 | 0 | [7 empty] | 0 | 0
0.196 | 0 | [7 empty] | 0 | 0
0.197 | 0 | [7 empty] | 0 | 0
0.198 | 0 | [7 empty] | 0 | 0
0.199 | 0 | [7 empty] | 0 | 0
0.2 | 0 | [7 empty] | 0 | 0
0.201 | 0 | [7 empty] | 0 | 0
0.202 | 0 | [7 empty] | 0 | 0
0.203 | 0 | [7 empty] | 0 | 0
0.204 | 0 | [7 empty] | 0 | 0
0.205 | 0 | [7 empty] | 0 | 0
0.206 | 0 | [7 empty] | 0 | 0
0.207 | 0 | [7 empty] | 0 | 0
0.208 | 0 | [7 empty] | 0 | 0
0.209 | 0 | [7 empty] | 0 | 0
0.21 | 0 | [7 empty] | 0 | 0
0.211 | 0 | [7 empty] | 0 | 0
0.212 | 0 | [7 empty] | 0 | 0
0.213 | 0 | [7 empty] | 0 | 0
0.214 | 0 | [7 empty] | 0 | 0
0.215 | 0 | [7 empty] | 0 | 0
0.216 | 0 | [7 empty] | 0 | 0
0.217 | 0 | [7 empty] | 0 | 0
0.218 | 0 | [7 empty] | 0 | 0
0.219 | 0 | [7 empty] | 0 | 0
0.22 | 0 | [7 empty] | 0 | 0
0.221 | 0 | [7 empty] | 0 | 0
0.222 | 0 | [7 empty] | 0 | 0
0.223 | 0 | [7 empty] | 0 | 0
0.224 | 0 | [7 empty] | 0 | 0
0.225 | 0 | [7 empty] | 0 | 0
0.226 | 0 | [7 empty] | 0 | 0
0.227 | 0 | [7 empty] | 0 | 0
0.228 | 0 | [6 empty] | The current supplied from pilot on both PT1000 and NTC thermal sensor shows a gaussian distribution in full agree with Giovanni studies. Only some Half-Staves show a different current values. | 0 | 0
0.229 | 0 | [7 empty] | 0 | 0
0.23 | 0 | [7 empty] | 0 | 0
0.231 | 0 | [7 empty] | 0 | 0
0.232 | 0 | [6 empty] | Statistical distribution of the temperatures error if assume a constant I APILOT current = mean value found. | 0 | 0
0.233 | 0 | [7 empty] | 0 | 0
0.234 | 0 | [7 empty] | 0 | 0
0.235 | 0 | [7 empty] | 0.0019243 | 0.00016036
0.236 | 1 | check on Sector 4 side C pos 1 | [6 empty] | 0.01652 | 0.00236
0.237 | 0 | [7 empty] | 0.11353 | 0.016219
0.238 | 0 | [7 empty] | 0.62459 | 0.089228
0.239 | 1 | check on Sector 10 side A pos 2 | [6 empty] | 2.7507 | 0.39296
0.24 | 0 | [7 empty] | 9.6975 | 1.3854
0.241 | 1 | [7 empty] | 27.3682 | 3.9097
0.242 | 4 | [7 empty] | 61.8305 | 8.8329
0.243 | 13 | [7 empty] | 111.8224 | 15.9746
0.244 | 18 | [7 empty] | 161.8918 | 23.1274
0.245 | 20 |  | 0.24516 | 0.00212 | 0.238 | 0.251 | 0.9% |  | 187.6249 | 26.8036
0.246 | 28 | [7 empty] | 174.0709 | 24.8673
0.247 | 23 | [7 empty] | 129.2801 | 18.4686
0.248 | 9 | [7 empty] | 76.8612 | 10.9802
0.249 | 6 | [7 empty] | 36.5807 | 5.2258
0.25 | 1 | [7 empty] | 13.9369 | 1.991
0.251 | 1 | [7 empty] | 4.2506 | 0.60723
0.252 | 2 | [7 empty] | 1.0378 | 0.14825
0.253 | 0 | [7 empty] | 0.20283 | 0.028976
0.254 | 0 | [7 empty] | 0.031734 | 0.0045334
0.255 | 0 | [7 empty] | 0.0039746 | 0.0005678
0.256 | 0 | [7 empty] | 0.0003985 | 5.6928e-05
0.257 | 0 | [7 empty] | 3.1984e-05 | 4.5691e-06
0.258 | 0 | [7 empty] | 0 | 0
0.259 | 0 | [7 empty] | 0 | 0
0.26 | 0 | [7 empty] | 0 | 0
0.261 | 1 | check on Sector 8 side A pos 5 | [6 empty] | 0 | 0
0.262 | 0 | [7 empty] | 0 | 0
0.263 | 0 | [7 empty] | 0 | 0
0.264 | 0 | [7 empty] | 0 | 0
0.265 | 0 | [7 empty] | 0 | 0
0.266 | 0 | [7 empty] | 0 | 0
0.267 | 0 | [7 empty] | 0 | 0
0.268 | 0 | [7 empty] | 0 | 0
0.269 | 0 | [7 empty] | 0 | 0
0.27 | 0 | [7 empty] | 0 | 0
0.271 | 0 | [7 empty] | 0 | 0
0.272 | 0 | [7 empty] | 0 | 0
0.273 | 0 | [7 empty] | 0 | 0
0.274 | 0 | [7 empty] | 0 | 0
0.275 | 0 | [7 empty] | 0 | 0
0.276 | 0 | [7 empty] | 0 | 0
0.277 | 0 | [7 empty] | 0 | 0
0.278 | 0 | [7 empty] | 0 | 0
0.279 | 0 | [7 empty] | 0 | 0
0.28 | 0 | [7 empty] | 0 | 0
0.281 | 1 | check on Sector 10side A pos 5 | [6 empty] | 0 | 0
0.282 | 0 | [7 empty] | 0 | 0
T MCM statistical
Class | Frequency | Comments | Mean | STD | Min | Max
13 | 0 |  | 20.0772 | 1.561 | 14.4467 | 26.1162
14 | 0
15 | 1 | check
16 | 0
17 | 0
18 | 8
19 | 21
20 | 36
21 | 31
22 | 20
23 | 7
24 | 4
25 | 0
26 | 0
27 | 1 | check
430 | 1 | 1AE
440 | 0 | 1B8
450 | 0 | 1C2
460 | 0 | 1CC
470 | 0 | 1D6
480 | 0 | 1E0
490 | 0 | 1EA
500 | 0 | 1F4
510 | 0 | 1FE
520 | 0 | 208
530 | 0 | 212
540 | 0 | 21C
550 | 0 | 226
560 | 0 | 230
570 | 0 | 23A
580 | 0 | 244
590 | 0 | 24E
600 | 0 | 258
610 | 0 | 262
620 | 1 | 26C
630 | 1 | 276
640 | 5 | 280
650 | 32 | 28A
660 | 56 | 294
670 | 27 | 29E
680 | 3 | 2A8
690 | 2 | 2B2
700 | 0 | 2BC
710 | 0 | 2C6
720 | 1 | 2D0
730 | 0 | 2DA
740 | 0 | 2E4
750 | 0 | 2EE
760 | 0 | 2F8
770 | 0 | 302
780 | 0 | 30C
790 | 0 | 316
800 | 0 | 320
810 | 1 | 32A
820 | 0 | 334
830 | 0 | 33E
840 | 0 | 348
850 | 0 | 352
860 | 0 | 35C
870 | 0 | 366
880 | 0 | 370
890 | 0 | 37A
900 | 0 | 384
910 | 0 | 38E
Altro | 0 | #VALUE!

## Error with Iap costant
Identification |  |  | Router TMCM (HEX) | TMCM Router [°C] | RMCM Router [Ω] | Router TBUS (HEX) | Thermal- camera TC [°C] | RBUS Router [Ω] | IAPILOT [mA] | TBUS Router [°C] | D T [°C]
Sector | Side | Half-Stave
0 | C | CH0 | 26E | 20.6518 | 5592.5707 | 267 | 28 | 5553.1824 | 0.2451 | 28.7367 | -0.73675
|  | CH1 | 2A0 | 19.4471 | 5874.022 | 25C | 26.5 | 5491.2933 |  | 25.5217 | 0.97827
|  | CH2 | 298 | 19.6353 | 5828.9898 | 253 | 25.5 | 5440.6569 |  | 22.8913 | 2.6087
|  | CH3 | 285 | 20.0891 | 5722.0383 | 268 | 26 | 5558.8086 |  | 29.029 | -3.029
|  | CH4 | 27F | 20.2344 | 5688.2642 | 267 | 28 | 5553.1824 |  | 28.7367 | -0.73675
|  | CH5 | 28D | 19.8968 | 5767.0705 | 265 | 26.5 | 5541.9298 |  | 28.1522 | -1.6522
| A | CH0 | 2BE | 18.7561 | 6042.8928 | 268 | 27 | 5558.8086 |  | 29.029 | -2.029
|  | CH1 | 2C8 | 18.5308 | 6099.183 | 262 | 28 | 5525.051 |  | 27.2754 | 0.72462
|  | CH2 | 2AD | 19.1449 | 5947.1993 | 264 | 27 | 5536.3035 |  | 27.8599 | -0.85992
|  | CH3 | [5 empty] | 2093.0233 |  |  | 0
|  | CH4 | 2A0 | 19.4471 | 5874.022 | 26B | 31 | 5575.6875 |  | 29.9058 | 1.0942
|  | CH5 | 2AC | 19.168 | 5941.5703 | 267 | 28 | 5553.1824 |  | 28.7367 | -0.73675
1 | C | CH0 | 269 | 20.7762 | 5564.4256 | 267 | 28 | 5553.1824 |  | 28.7367 | -0.73675
|  | CH1 | 290 | 19.8251 | 5783.9576 | 258 | 26 | 5468.7882 |  | 24.3526 | 1.6474
|  | CH2 | 29E | 19.494 | 5862.764 | 252 | 27 | 5435.0306 |  | 22.599 | 4.401
|  | CH3 | 2AB | 19.1911 | 5935.9413 | 262 | 25 | 5525.051 |  | 27.2754 | -2.2754
|  | CH4 | 291 | 19.8013 | 5789.5866 | 265 | 28 | 5541.9298 |  | 28.1522 | -0.1522
|  | CH5 | 292 | 19.7775 | 5795.2157 | 264 | 28 | 5536.3035 |  | 27.8599 | 0.14008
| A | CH0 | 2CD | 18.4191 | 6127.3282 | 266 | 28 | 5547.5561 |  | 28.4445 | -0.44447
|  | CH1 | 2BE | 18.7561 | 6042.8928 | 264 | 29 | 5536.3035 |  | 27.8599 | 1.1401
|  | CH2 | 29B | 19.5645 | 5845.8769 | 25E | 25 | 5502.5459 |  | 26.1063 | -1.1063
|  | CH3 | 277 | 20.4298 | 5643.232 | 25D | 27 | 5496.9196 |  | 25.814 | 1.186
|  | CH4 | 29F | 19.4705 | 5868.393 | 259 | 26 | 5474.4145 |  | 24.6449 | 1.3551
|  | CH5 | 2AC | 19.168 | 5941.5703 | 265 | 28 | 5541.9298 |  | 28.1522 | -0.1522
2 | C | CH0 | 240 | 21.8245 | 5333.6355 | 268 | 29.1 | 5558.8086 |  | 29.029 | 0.070979
|  | CH1 | 263 | 20.9264 | 5530.6514 | 268 | 29.5 | 5558.8086 |  | 29.029 | 0.47098
|  | CH2 | 277 | 20.4298 | 5643.232 | 277 | 28.9 | 5643.2028 |  | 33.4131 | -4.5131
|  | CH3 | 24E | 21.4607 | 5412.4419 | 276 | 28.9 | 5637.5765 |  | 33.1209 | -4.2209
|  | CH4 | 268 | 20.8012 | 5558.7966 | 26B | 27.7 | 5575.6875 |  | 29.9058 | -2.2058
|  | CH5 | 25C | 21.103 | 5491.2483 | 263 | 28 | 5530.6773 |  | 27.5677 | 0.43235
| A | CH0 | 29B | 19.5645 | 5845.8769 | 27A | 30.4 | 5660.0816 |  | 34.29 | -3.89
|  | CH1 | 28C | 19.9207 | 5761.4415 | 26A | 29.6 | 5570.0612 |  | 29.6136 | -0.013569
|  | CH2 | 23F | 21.8507 | 5328.0065 | 274 | 28.3 | 5626.3239 |  | 32.5363 | -4.2363
|  | CH3 | 228 | 22.463 | 5198.5389 | 26B | 30 | 5575.6875 |  | 29.9058 | 0.094157
|  | CH4 | 231 | 22.2213 | 5249.2002 | 25F | 27.9 | 5508.1722 |  | 26.3986 | 1.5014
|  | CH5 | 247 | 21.6418 | 5373.0387 | 265 | 26 | 5541.9298 |  | 28.1522 | -2.1522
3 | C | CH0 | 29F | 19.4705 | 5868.393 | 268 | 28.7 | 5558.8086 |  | 29.029 | -0.32902
|  | CH1 | 287 | 20.0408 | 5733.2964 | 272 | 33.2 | 5615.0714 |  | 31.9518 | 1.2482
|  | CH2 | 26F | 20.6271 | 5598.1998 | 271 | 30.1 | 5609.4451 |  | 31.6595 | -1.5595
|  | CH3 | 255 | 21.2811 | 5451.8451 | 273 | 28.7 | 5620.6977 |  | 32.244 | -3.544
|  | CH4 | 1AD | 26.0716 | 4506.1688 | 267 | 28.2 | 5553.1824 |  | 28.7367 | -0.53675
|  | CH5 | 267 | 20.8261 | 5553.1675 | 25F | 27.8 | 5508.1722 |  | 26.3986 | 1.4014
| A | CH0 | 2BA | 18.8469 | 6020.3767 | 264 | 27.2 | 5536.3035 |  | 27.8599 | -0.65992
|  | CH1 | [5 empty] | 2093.0233 |  |  | 0
|  | CH2 | 290 | 19.8251 | 5783.9576 | 270 | 29.6 | 5603.8188 |  | 31.3672 | -1.7672
|  | CH3 | 25D | 21.0777 | 5496.8773 | 26B | 28 | 5575.6875 |  | 29.9058 | -1.9058
|  | CH4 | 2B0 | 19.0757 | 5964.0864 | 261 | 27.1 | 5519.4247 |  | 26.9831 | 0.1169
|  | CH5 | 28A | 19.9687 | 5750.1834 | 265 | 28.9 | 5541.9298 |  | 28.1522 | 0.7478
4 | C | CH0 | 276 | 20.4544 | 5637.6029 | 26C | 29.1 | 5581.3137 |  | 30.1981 | -1.0981
|  | CH1 | [5 empty] | 2093.0233
|  | CH2 | 258 | 21.2046 | 5468.7322 | 269 | 29.1 | 5564.4349 |  | 29.3213 | -0.22129
|  | CH3 | 210 | 23.1212 | 5063.4423 | 26F | 30.3 | 5598.1926 |  | 31.0749 | -0.77494
|  | CH4 | [5 empty] | 2093.0233
|  | CH5 | 266 | 20.8512 | 5547.5385 | 266 | 29.8 | 5547.5561 |  | 28.4445 | 1.3555
| A | CH0 | 2DD | 18.0654 | 6217.3926 | 26A | 27.4 | 5570.0612 |  | 29.6136 | -2.2136
|  | CH1 | hot !!! |  |  |  |  | 2093.0233
|  | CH2 | 27C | 20.3075 | 5671.3771 | 25F | 28.1 | 5508.1722 |  | 26.3986 | 1.7014
|  | CH3 | 209 | 23.3171 | 5024.0391 | 266 | 30.3 | 5547.5561 |  | 28.4445 | 1.8555
|  | CH4 | 224 | 22.5713 | 5176.0228 | 285 | 30.9 | 5721.9706 |  | 37.505 | -6.605
|  | CH5 | 258 | 21.2046 | 5468.7322 | 25E | 29.3 | 5502.5459 |  | 26.1063 | 3.1937
5 | C | CH0 | 266 | 20.8512 | 5547.5385 | 265 | 31.4 | 5541.9298 |  | 28.1522 | 3.2478
|  | CH1 | 26D | 20.6767 | 5586.9417 | 26A | 30.4 | 5570.0612 |  | 29.6136 | 0.78643
|  | CH2 | 260 | 21.0019 | 5513.7644 | 25E | 28.2 | 5502.5459 |  | 26.1063 | 2.0937
|  | CH3 | 23F | 21.8507 | 5328.0065 | 26E | 28.9 | 5592.5663 |  | 30.7827 | -1.8827
|  | CH4 | 23D | 21.9032 | 5316.7485 | 26F | 30 | 5598.1926 |  | 31.0749 | -1.0749
|  | CH5 | 24E | 21.4607 | 5412.4419 | 268 | 31.6 | 5558.8086 |  | 29.029 | 2.571
| A | CH0 | 2B0 | 19.0757 | 5964.0864 | 263 | 29.2 | 5530.6773 |  | 27.5677 | 1.6323
|  | CH1 | 2CA | 18.486 | 6110.4411 | 260 | 28.7 | 5513.7984 |  | 26.6908 | 2.0092
|  | CH2 | 26F | 20.6271 | 5598.1998 | 26D | 28.9 | 5586.94 |  | 30.4904 | -1.5904
|  | CH3 | 26D | 20.6767 | 5586.9417 | 272 | 29.3 | 5615.0714 |  | 31.9518 | -2.6518
|  | CH4 | 272 | 20.5529 | 5615.0868 | 25E | 30.2 | 5502.5459 |  | 26.1063 | 4.0937
|  | CH5 | 280 | 20.2101 | 5693.8932 | 263 | 29.5 | 5530.6773 |  | 27.5677 | 1.9323
6 | C | CH0 | 22C | 22.3552 | 5221.055 | 267 | 28.7 | 5553.1824 |  | 28.7367 | -0.036747
|  | CH1 | 23D | 21.9032 | 5316.7485 | 261 | 29.1 | 5519.4247 |  | 26.9831 | 2.1169
|  | CH2 | 25C | 21.103 | 5491.2483 | 261 | 26.4 | 5519.4247 |  | 26.9831 | -0.5831
|  | CH3 | 25C | 21.103 | 5491.2483 | 26E | 29.1 | 5592.5663 |  | 30.7827 | -1.6827
|  | CH4 | 21B | 22.817 | 5125.3616 | 271 | 29.7 | 5609.4451 |  | 31.6595 | -1.9595
|  | CH5 | 251 | 21.3835 | 5429.329 | 26D | 28.5 | 5586.94 |  | 30.4904 | -1.9904
| A | CH0 | 283 | 20.1374 | 5710.7803 | 253 | 25.5 | 5440.6569 |  | 22.8913 | 2.6087
|  | CH1 | 293 | 19.7537 | 5800.8447 | 25E | 26.8 | 5502.5459 |  | 26.1063 | 0.69372
|  | CH2 | 263 | 20.9264 | 5530.6514 | 25C | 25.5 | 5491.2933 |  | 25.5217 | -0.021732
|  | CH3 | 27D | 20.2831 | 5677.0061 | 26A | 26.3 | 5570.0612 |  | 29.6136 | -3.3136
|  | CH4 | 244 | 21.7199 | 5356.1517 | 256 | 26.4 | 5457.5357 |  | 23.7681 | 2.6319
|  | CH5 | 267 | 20.8261 | 5553.1675 | 25B | 28 | 5485.6671 |  | 25.2295 | 2.7705
7 | C | CH0 | 27D | 20.2831 | 5677.0061 | 263 | 27.5 | 5530.6773 |  | 27.5677 | -0.067651
|  | CH1 | 2A3 | 19.377 | 5890.9091 | 260 | 28 | 5513.7984 |  | 26.6908 | 1.3092
|  | CH2 | 28F | 19.849 | 5778.3286 | 26A | 28.5 | 5570.0612 |  | 29.6136 | -1.1136
|  | CH3 | 28C | 19.9207 | 5761.4415 | 260 | 27.5 | 5513.7984 |  | 26.6908 | 0.80917
|  | CH4 | 278 | 20.4053 | 5648.861 | 253 | 26.5 | 5440.6569 |  | 22.8913 | 3.6087
|  | CH5 | 25A | 21.1537 | 5479.9902 | 265 | 30 | 5541.9298 |  | 28.1522 | 1.8478
| A | CH0 | 2FA | 17.4394 | 6380.6343 | 264 | 25.5 | 5536.3035 |  | 27.8599 | -2.3599
|  | CH1 | 2EB | 17.7609 | 6296.1989 | 267 | 29 | 5553.1824 |  | 28.7367 | 0.26325
|  | CH2 | 2B4 | 18.9839 | 5986.6025 | 26F | 30 | 5598.1926 |  | 31.0749 | -1.0749
|  | CH3 | 294 | 19.73 | 5806.4737 | 268 | 29.6 | 5558.8086 |  | 29.029 | 0.57098
|  | CH4 | 2AF | 19.0988 | 5958.4574 | 264 | 30 | 5536.3035 |  | 27.8599 | 2.1401
|  | CH5 | 297 | 19.6589 | 5823.3608 | 25F | 29 | 5508.1722 |  | 26.3986 | 2.6014
8 | C | CH0 | 290 | 19.8251 | 5783.9576 | 267 | 28.1 | 5553.1824 |  | 28.7367 | -0.63675
|  | CH1 | 262 | 20.9515 | 5525.0224 | 266 | 28 | 5547.5561 |  | 28.4445 | -0.44447
|  | CH2 | 27B | 20.3319 | 5665.7481 | 266 | 23.8 | 5547.5561 |  | 28.4445 | -4.6445
|  | CH3 | 2A1 | 19.4237 | 5879.651 | 258 | 27.4 | 5468.7882 |  | 24.3526 | 3.0474
|  | CH4 | 263 | 20.9264 | 5530.6514 | 26B | 28.7 | 5575.6875 |  | 29.9058 | -1.2058
|  | CH5 | 254 | 21.3066 | 5446.2161 | 261 | 28.5 | 5519.4247 |  | 26.9831 | 1.5169
| A | CH0 | 31B | 16.7494 | 6566.3922 | 268 | 25.5 | 5558.8086 |  | 29.029 | -3.529
|  | CH1 | 2FF | 17.3334 | 6408.7795 | 272 | 26 | 5615.0714 |  | 31.9518 | -5.9518
|  | CH2 | 2B6 | 18.9382 | 5997.8606 | 254 | 26.1 | 5446.2831 |  | 23.1835 | 2.9165
|  | CH3 | 2D0 | 18.3523 | 6144.2152 | 25D | 25.3 | 5496.9196 |  | 25.814 | -0.51401
|  | CH4 | 2AE | 19.1218 | 5952.8284 | 258 | 26.7 | 5468.7882 |  | 24.3526 | 2.3474
|  | CH5 | [5 empty] | 2093.0233
9 | C | CH0 | 2CE | 18.3968 | 6132.9572 | 248 | 23.8 | 5378.7678 |  | 19.6763 | 4.1237
|  | CH1 | 2CF | 18.3745 | 6138.5862 | 26C | 29.6 | 5581.3137 |  | 30.1981 | -0.59812
|  | CH2 | 2A6 | 19.3071 | 5907.7962 | 260 | 26.6 | 5513.7984 |  | 26.6908 | -0.090829
|  | CH3 | 25D | 21.0777 | 5496.8773 | 268 | 28.7 | 5558.8086 |  | 29.029 | -0.32902
|  | CH4 | 2A6 | 19.3071 | 5907.7962 | 255 | 27 | 5451.9094 |  | 23.4758 | 3.5242
|  | CH5 | 2B8 | 18.8925 | 6009.1186 | 259 | 29.3 | 5474.4145 |  | 24.6449 | 4.6551
| A | CH0 | 2F3 | 17.5888 | 6341.2311 | 263 | 24.6 | 5530.6773 |  | 27.5677 | -2.9677
|  | CH1 | 316 | 16.8525 | 6538.247 | 26B | 26.7 | 5575.6875 |  | 29.9058 | -3.2058
|  | CH2 | 2DF | 18.0216 | 6228.6506 | 258 | 26.6 | 5468.7882 |  | 24.3526 | 2.2474
|  | CH3 | 262 | 20.9515 | 5525.0224 | 263 | 26.6 | 5530.6773 |  | 27.5677 | -0.96765
|  | CH4 | 2E0 | 17.9998 | 6234.2797 | 263 | 26.4 | 5530.6773 |  | 27.5677 | -1.1677
|  | CH5 | 293 | 19.7537 | 5800.8447 | 269 | 28.2 | 5564.4349 |  | 29.3213 | -1.1213
10 | C | CH0 | 2E2 | 17.9561 | 6245.5377 | 26A | 26.7 | 5570.0612 |  | 29.6136 | -2.9136
|  | CH1 | 2A7 | 19.2838 | 5913.4252 | 25C | 28.4 | 5491.2933 |  | 25.5217 | 2.8783
|  | CH2 | 2C1 | 18.6883 | 6059.7799 | 267 | 29 | 5553.1824 |  | 28.7367 | 0.26325
|  | CH3 | 2D5 | 18.2415 | 6172.3604 | 26A | 27.2 | 5570.0612 |  | 29.6136 | -2.4136
|  | CH4 | 2D8 | 18.1753 | 6189.2475 | 262 | 28.1 | 5525.051 |  | 27.2754 | 0.82462
|  | CH5 | 25B | 21.1283 | 5485.6192 | 252 | 27.6 | 5435.0306 |  | 22.599 | 5.001
| A | CH0 | 301 | 17.2911 | 6420.0375 | 261 | 27.4 | 5519.4247 |  | 26.9831 | 0.4169
|  | CH1 | 24F | 21.4349 | 5418.0709 | 26F | 29 | 5598.1926 |  | 31.0749 | -2.0749
|  | CH2 | 2A0 | 19.4471 | 5874.022 | 23E | 24.9 | 5322.5051 |  | 16.7535 | 8.1465
|  | CH3 | 2A9 | 19.2374 | 5924.6832 | 25C | 26.6 | 5491.2933 |  | 25.5217 | 1.0783
|  | CH4 | 25B | 21.1283 | 5485.6192 | 264 | 28.3 | 5536.3035 |  | 27.8599 | 0.44008
CH5
Error on temperatures calculation if IAPILOT is kept with a costant value = at mean distrubution value=0,2454
Class | Frequency | Comments | Mean | STD | Min | Max
-7 | 0 |  | 0.015414 | 2.3371 | -6.605 | 8.1465
-6 | 1
-5 | 1
-4 | 4
-3 | 6
-2 | 11
-1 | 17
0 | 27
1 | 18
2 | 17
3 | 14
4 | 5
5 | 4
6 | 1
7 | 0
8 | 0
9 | 1

## Measurements done in Bari
HALF_STAVE | MCM | Status | Production_Date | T Read SW (MCM) | T Read SW (Pixel Bus) | T Read HW_Ext_Chain (Pixel Bus) | T chuck | DT (Pixel Bus) | DT (MCM) |  | Temp | Frequency MCM | Frequency Bus
HS_001 | . | NOT | 2003-12-01 | [7 empty] | 15 | 0
HS_002 | . | NOT | 2004-02-18 | [7 empty] | 16 | 1
HS_003 | Cerami | NOT | 2004-03-12 | [7 empty] | 17 | 0
HS_004 | Cerami | NOT | 2004-06-25 | [7 empty] | 18 | 6
HS_005 | PCB du | NOT | 2005-01-10 | [7 empty] | 19 | 23
HS_006 | PCB du | NOT | 2005-01-10 | [7 empty] | 20 | 19
HS_007R | RRX001 | NOT | 2005-01-28 | [7 empty] | 21 | 16
HS_008L | LRX004 | NOT | 2005-02-16 | [7 empty] | 22 | 18
HS_009R | RRX005 | Sector 0 | 2005-02-21 | [7 empty] | 23 | 16
HS_010R | RRX002 | Sector 0 | 2005-02-28 | [7 empty] | 24 | 9 | 0
HS_011L | LRX001 | NOT | 2005-03-10 | [7 empty] | 25 | 5 | 3
HS_012L | LRX002 | NOT | 2005-03-15 | [7 empty] | 26 | 2 | 9
HS_013L | LRX005 | Sector 0 | 2005-04-20 | [7 empty] | 27 | 2 | 8
HS_014R | RRX003 | For WB Test | 2005-04-26 | [7 empty] | 28 | 0 | 21
HS_015L | LRX003 | Sector 0 | 2005-04-28 | [7 empty] | 29 | 1 | 15
HS_016R | R_001 | Sector 0 | 2005-05-03 | [7 empty] | 30 |  | 12
HS_017L | L_001 | NOT | 2005-05-27 | [7 empty] | 31 |  | 8
HS_018L | L_004 | Sector 0 | 2005-06-06 | [7 empty] | 32 |  | 11
HS_019R | R_004 | Sector 0 | 2005-06-06 | [7 empty] | 33 |  | 10
HS_020L | L_002 | NOT | 2005-06-07 | [7 empty] | 34 |  | 6
HS_021R | R_002 | Dis-Assembled | 2005-06-16 | [7 empty] | 35 |  | 7
HS_022L | L_005 | Sector 0 | 2005-07-12 | [7 empty] | 36 |  | 5
HS_023L | L_009 | Sector 0 | 2005-07-28 | [7 empty] | 37 |  | 2
HS_024R | R_013 | Sector 0 | 2005-07-25 | [7 empty] | 38 |  | 3
HS_025R | R_010 | Sector 0 | 2005-09-09
HS_026L | L_011 | Sector 0 | 2005-09-09
HS_027R | R_007 | Sector 1 | 2005-12-13
HS_028L | LP09 | Sector 1 | 2005-12-13
HS_029L | LP08 | Dis-Assembled | 2005-11-22
HS_030R | RP06 | Sector 1 | 2005-12-01
HS_031L | LP10 | Sector 1 | 2005-12-01
HS_032R | RP07 | Sector 1 | 2005-12-20 | 22 |  | 29
HS_033R | RP04 | Sector 1 | 2006-01-04 | 19 |  | 27 | 15.2 |  | 3.8
HS_034L | LP02 | Sector 1 | 2005-12-20 | 20 |  | 25.5
HS_035L | LP06 | BAD | 2006-01-04
HS_036R | RP05 | Sector 1 | 2006-01-03 | 23 |  | 27 | 15.2 |  | 7.8
HS_037L | LP07 | Sector 1 | 2006-01-02 | 20.32 |  | 30 | 15.6 |  | 4.72
HS_038R | RP01 | NOT | 2006-01-13
HS_039L | LP05 | Sector 1 | 2006-01-13 | 19.43 |  | 30.8 | 15.9 |  | 3.53
HS_040R | R_006 | Sector 2 | 2006-01-31 | 20 |  | 30 | 16 |  | 4
HS_041R | R08 | NOT | 2006-02-01
HS_042L | L_010 | Sector 2 | 2006-01-31 | 20.3 |  | 28.6 | 15.2 |  | 5.1
HS_043L | LP03 | Sector 2 | 2006-03-28 | 23.5 |  | 35.8 | 20.4 |  | 3.1
HS_044R | RP03 | Sector 2 | 2006-04-05 | 22.2 |  | 34.5 | 17.3 |  | 4.9
HS_045R | RP11 | Sector 2 | 2006-03-28 | 21.3 |  | 32 | 22.6 |  | -1.3
HS_046L | L_014 | Sector 2 | 2006-03-22 | 22 |  | 31 | 21 |  | 1
HS_047R | RP12 | Sector 2 | 2006-03-10 | 23 |  | 31 | 20 |  | 3
HS_048L | L_015 | Sector 2 | 2006-03-10 | 24 |  | 34 | 20 |  | 4
HS_049R | RP09 | Sector 2 | 2006-03-22 | 22.4 |  | 31 | 17.2 |  | 5.2
HS_050R | R_008 | Sector 2 | 2006-04-21 | 21.4 |  | 37.6 | 18.8 |  | 2.6
HS_051L | L_012 | Sector 2 | 2006-04-05 | 22.9 |  | 36.5 | 22.8 |  | 0.1
HS_052L | LP20 | Dis-Assembled | 2006-04-11
HS_053R | RP13 | Under check | 2006-04-20
HS_054L | LP11 | Sector 4 | 2006-07-06 | 21 |  | 37.5 | 18.7 |  | 2.3
HS_055L | LP04 | Sector 2 | 2006-04-21 | 23.8 |  | 35 | 20 |  | 3.8
HS_056L | LP01 | Sector 3 | 2006-05-04 | 24 |  | 33 | 22 |  | 2
HS_057R | RP02 | Sector 3 | 2006-05-04 | 25 |  | 35.5 | 20 |  | 5
HS_058R | RP15 | Sector 10 | 2007-01-29 | 20 | 35 | 37.4 | 18.7 | -2.4 | 1.3
HS_059L | LP15 | Tested: OK | 2007-02-19 | 22 | 34 | 31.4 | 18.1 | 2.6 | 3.9
HS_060R | RP10 | Sector 3 | 2006-06-01 | 21.2 |  | 33.5 | 22.3 |  | -1.1
HS_061L | LP24 | Sector 3 | 2006-06-01 | 20.7 |  | 29.5 | 18.2 |  | 2.5
HS_062L | LP22 | Sector 3 | 2006-06-07 | 24.3 |  | 31.1 | 18.8 |  | 5.5
HS_063R | RP19 | Sector 3 | 2006-06-07 |  |  | 30.7 | 18.9
HS_064R | RP20 | Sector 3 | 2006-06-15 | 24.7 |  | 29 | 18.7 |  | 6
HS_065L | LP17 | Sector 3 | 2006-06-15 | 21.3 |  | 31 | 18.2 |  | 3.1
HS_066R | RP14 | Sector 7 | 2006-10-25 | 19.8 |  | 30 | 20.5 |  | -0.7
HS_067L | LP21 | Sector 3 | 2006-06-21 | 22 |  | 28.3 | 18.1 |  | 3.9
HS_068R | RP18 | Sector 3 | 2006-06-21 | 20 |  | 30 | 20 |  | 0
HS_069L | LP23 | Sector 3 | 2006-06-30 | 21.7 |  | 31.8 | 19.5 |  | 2.2
HS_070R | RP17 | Sector 3 | 2006-06-30 | 26 |  | 35 | 19.5 |  | 6.5
HS_071R | RP16 | Sector 4 | 2006-07-03 | 23 |  | 31.5 | 17.6 |  | 5.4
HS_072L | LP16 | Sector 4 | 2006-07-04 | 23.8 |  | 32.5 | 19.5 |  | 4.3
HS_073L | LP28 | Sector 4 | 2006-07-14 | 25 |  | 36 | 19 |  | 6
HS_074R | RP26 | Sector 4 | 2006-07-06 | 23 |  | 29 | 18 |  | 5
HS_075R | RP30 | Dis-Assembled
HS_076L | LP27 | Sector 4 | 2006-07-19 | 22.9 |  | 32.5 | 19.8 |  | 3.1
HS_077L | L08 | Sector 4 | 2006-07-26 |  |  | 33 | 18
HS_078R | R09 | Sector 4 | 2006-07-14 | 22.6 |  | 29.7 | 19.4 |  | 3.2
HS_079R | RP24 | Sector 4 | 2006-07-19 | 22.3 |  | 25.6 | 18.8 |  | 3.5
HS_080R | RP21 | Sector 4 | 2006-07-26 | 22 |  | 31 | 19 |  | 3
HS_081L | LP13 | Sector 4 | 2006-07-28 | 23.5 |  | 34.3 | 18.9 |  | 4.6
HS_082R | RP25 | Sector 4 | 2006-07-28 | 23 |  | 33 | 20 |  | 3
HS_083L | LP31 | Sector 1 | 2006-10-10 | 29 |  | 36 | 20 |  | 9
HS_084R | RP29 | Sector 6 | 2006-10-16 | 23 |  | 33.5 | 19.5 |  | 3.5
HS_085L | LP18 | Sector 5 | 2006-09-11 | 23 |  | 33 | 19 |  | 4
HS_086R | RP32 | Under check | 2006-07-31
HS_087L | LP33 | Sector 5 | 2006-08-18 | 22 |  | 35 | 19 |  | 3
HS_088R | RP35 | Sector 5 | 2006-08-18 | 22 |  | 32 | 20 |  | 2
HS_089L | LP36 | Sector 5 | 2006-08-18 | 24 |  | 36.4 | 18.9 |  | 5.1
HS_090R | RP34 | Sector 5 | 2006-08-18 | 24 |  | 32 | 18.6 |  | 5.4
HS_091R | RP38 | Sector 6 | 2006-09-01 | 23 |  | 31 | 19 |  | 4
HS_092L | LP35 | Sector 6 | 2006-08-31 | 22 |  | 35 | 19 |  | 3
HS_093R | RP47 | Sector 5 | 2006-09-07 | 19 |  | 26 | 19 |  | 0
HS_094L | LP25 | Sector 5 | 2006-09-07 | 27 |  | 32 | 19 |  | 8
HS_095R | RP40 | Sector 5 | 2006-09-11 | 21 |  | 33 | 20 |  | 1
HS_096L | LP29 | Sector 8 | 2006-11-09 | 19 |  | 30 | 18 |  | 1
HS_097R | RP49 | Sector 5 | 2006-09-15 | 21 |  | 33 | 19 |  | 2
HS_098L | LP26 | Sector 5 | 2006-09-15 | 23.7 |  | 31.2 | 19.6 |  | 4.1
HS_099R | RP44 | Sector 5 | 2006-09-20 | 16 | 25 | 28.5 | 17 | -3.5 | -1 | 5481.25 | 1288.0938 | 254
HS_100L | LP30 | Sector 5 | 2006-09-20 | 20 | 29 | 32.7 | 18.5 | -3.7 | 1.5 | 5558.25 | 1306.1888 | 262
HS_101R | RP45 | Sector 6 | 2006-09-25 | 20.6 | 31.2 | 34.3 | 19.7 | -3.1 | 0.9 | 5600.6 | 1316.141 | 269
HS_102L | LP34 | Sector 6 | 2006-09-25 | 21.2 | 25.8 | 31.2 | 18.3 | -5.4 | 2.9 | 5496.65 | 1291.7127 | 257
HS_103R | RP46 | Sector 8 | 2006-10-09 | 25 | 32.6 | 33 | 24 | -0.4 | 1 | 5627.55 | 1322.4742 | 26E
HS_104L | LP48 | Sector 6 | 2006-09-26 | 25.4 | 29.8 | 35.6 | 23.7 | -5.8 | 1.7 | 5573.65 | 1309.8077 | 264
HS_105L | LP41 | Sector 6 | 2006-09-28 | 19.3 | 32 | 33.6 | 17.5 | -1.6 | 1.8 | 5616 | 1319.76 | 26C
HS_106R | RP41 | Sector 6 | 2006-10-04 | 18 | 24 | 28 | 18 | -4 | 0 | 5462 | 1283.57 | 250
HS_107L | LP40 | Sector 6 | 2006-10-04 | 22 | 30 | 34 | 20 | -4 | 2 | 5577.5 | 1310.7125 | 265
HS_108R | RP53 | Sector 6 | 2006-10-04 | 20 | 29 | 30 | 19.8 | -1 | 0.2 | 5558.25 | 1306.1888 | 262
HS_109L | LP37 | Sector 6 | 2006-10-13 | 23 | 31 | 32 | 22.6 | -1 | 0.4 | 5596.75 | 1315.2362 | 269
HS_110R | RP39 | Sector 6 | 2006-10-13 | 21 | 23 | 26 | 18.5 | -3 | 2.5 | 5442.75 | 1279.0462 | 24D
HS_111L | LP38 | Sector 7 | 2006-10-12 | 23 | 31.5 | 33.6 | 22 | -2.1 | 1 | 5606.375 | 1317.4981 | 26A
HS_112R | RP50 | Sector 1 | 2006-10-18 | 21 | 29 | 30 | 19 | -1 | 2 | 5558.25 | 1306.1888 | 262
HS_113R | RP55 | Dis-Assembled | 2006-10-24
HS_114L | LP50 | Sector 7 | 2006-10-20 | 20 | 28 | 28 | 18.7 | 0 | 1.3 | 5539 | 1301.665 | 25E
HS_115L | LP55 | Sector 7 | 2006-10-25 | 21 | 29 | 28.5 | 18.3 | 0.5 | 2.7 | 5558.25 | 1306.1888 | 262
HS_116R | RP42 | Sector 7 | 2006-10-30 | 19 | 29 | 28 | 18.5 | 1 | 0.5 | 5558.25 | 1306.1888 | 262
HS_117L | LP52 | Sector 7 | 2006-11-02 | 20 | 30 | 29 | 18.3 | 1 | 1.7 | 5577.5 | 1310.7125 | 265
HS_118R | RP58 | Sector 7 | 2006-10-31 | 19 | 25 | 25.7 | 18.5 | -0.7 | 0.5 | 5481.25 | 1288.0938 | 254
HS_119R | RP56 | Sector 7 | 2006-11-03 | 19 | 22 | 26.5 | 18 | -4.5 | 1 | 5423.5 | 1274.5225 | 249
HS_120L | LP53 | Sector 7 | 2006-11-06 | 20 | 29 | 27.8 | 18.4 | 1.2 | 1.6 | 5558.25 | 1306.1888 | 262
HS_121R | RP57 | Sector 7 | 2006-11-06 | 18 | 26 | 27.8 | 18.2 | -1.8 | -0.2 | 5500.5 | 1292.6175 | 257
HS_122R | RP59 | Sector 7 | 2006-11-08 | 19 | 28 | 27 | 18.5 | 1 | 0.5 | 5539 | 1301.665 | 25E
HS_123L | LP54 | Sector 7 | 2006-11-08 | 18 | 32 | 28.5 | 18.5 | 3.5 | -0.5 | 5616 | 1319.76 | 26C
HS_124R | RP52 | Sector 8 | 2006-11-10 | 18 | 29 | 28.6 | 18.7 | 0.4 | -0.7 | 5558.25 | 1306.1888 | 262
HS_125R | RP48 | Sector 8 | 2006-11-16 | 20 | 25 | 26.6 | 18.4 | -1.6 | 1.6 | 5481.25 | 1288.0938 | 254
HS_126L | LP43 | Sector 8 | 2006-11-16 | 21 | 29 | 28.4 | 18.4 | 0.6 | 2.6 | 5558.25 | 1306.1888 | 262
HS_127R | RP54 | Under check | 2006-11-20
HS_128R | RP43 | Sector 8 | 2006-11-30 | 22 | 27 | 27.2 | 18.6 | -0.2 | 3.4 | 5519.75 | 1297.1412 | 25B
HS_129L | LP39 | Sector 8 | 2006-11-17 | 21 | 28 | 28.1 | 18.1 | -0.1 | 2.9 | 5539 | 1301.665 | 25E
HS_130L | LP47 | Sector 8 | 2006-12-01 | 27 | 31 | 29.3 | 18.4 | 1.7 | 8.6 | 5596.75 | 1315.2362 | 269
HS_131L | LP66 | Sector 8 | 2006-11-29 | 20 | 27 | 27.5 | 18.5 | -0.5 | 1.5 | 5519.75 | 1297.1412 | 25B
HS_132R | RP64 | Sector 8 | 2006-12-01 | 19 | 24 | 25.3 | 18.4 | -1.3 | 0.6 | 5462 | 1283.57 | 250
HS_133L | LP56 | Dis-Assembled | 2006-12-04
HS_134R | RP60 | Sector 8 | 2006-12-05 | 20 | 28 | 27.5 | 18.5 | 0.5 | 1.5 | 5539 | 1301.665 | 25E
HS_135L | LP67 | Sector 8 | 2006-12-06 | 21 | 26 | 30 | 18.5 | -4 | 2.5 | 5500.5 | 1292.6175 | 257
HS_136L | LP59 | Sector 9 | 2006-12-11 | 21.5 | 27.5 | 27.5 | 18.5 | 0 | 3 | 5529.375 | 1299.4031 | 25C
HS_137R | RP65 | Sector 9 | 2006-12-11 | 20 | 22 | 24.2 | 18.4 | -2.2 | 1.6 | 5423.5 | 1274.5225 | 249
HS_138R | RP66 | Sector 9 | 2006-12-13 | 19 | 23 | 25 | 18.4 | -2 | 0.6 | 5442.75 | 1279.0462 | 24D
HS_139L | LP60 | Sector 9 | 2006-12-14 | 19 | 28 | 28 | 18.3 | 0 | 0.7 | 5539 | 1301.665 | 25E
HS_140R | RP63 | Sector 9 | 2006-12-19 | 18 | 25 | 25.6 | 18.3 | -0.6 | -0.3 | 5481.25 | 1288.0938 | 254
HS_141R | RP62 | BAD | 2006-12-28
HS_142L | LP63 | Sector 9 | 2006-12-19 | 19 | 25 | 27.5 | 18.3 | -2.5 | 0.7 | 5481.25 | 1288.0938 | 254
HS_143L | LP62 | Sector 9 | 2006-12-20 | 20 | 27 | 27.5 | 18.3 | -0.5 | 1.7 | 5519.75 | 1297.1412 | 25B
HS_144R | RP51 | NOT USED | 2007-01-04 |  |  |  |  | 0 | 0 | 5000 | 1175 | 1FD
HS_145L | LP44 | Sector 9 | 2007-01-09 | 19 | 28 | 26.1 | 18.5 | 1.9 | 0.5 | 5539 | 1301.665 | 25E
HS_146R | RP37 | Sector 9 | 2007-02-01 | 19 | 27 | 27.6 | 18.4 | -0.6 | 0.6 | 5519.75 | 1297.1412 | 25B
HS_147L | LP46 | Sector 9 | 2007-01-10 | 19 | 26 | 27.4 | 18.4 | -1.4 | 0.6 | 5500.5 | 1292.6175 | 257
HS_148R | RP70 | Sector 9 | 2007-01-10 | 19 | 28 | 27.5 | 18.2 | 0.5 | 0.8 | 5539 | 1301.665 | 25E
HS_149R | RP69 | Sector 9 | 2007-01-15 | 20 | 27 | 24.5 | 18.5 | 2.5 | 1.5 | 5519.75 | 1297.1412 | 25B
HS_150L | LP70 | Sector 10 | 2007-01-17 | 21 | 24 | 28.4 | 18.3 | -4.4 | 2.7 | 5462 | 1283.57 | 250
HS_151R | RP61 | Sector 10 | 2007-01-16 | 21.5 | 22.4 | 29.1 | 18.4 | -6.7 | 3.1 | 5431.2 | 1276.332 | 24B
HS_152R | RP67 | Sector 10 | 2007-01-19 | 18.1 | 26.7 | 26.5 | 18 | 0.2 | 0.1 | 5513.975 | 1295.7841 | 25A
HS_153L | LP42 | Sector 10 | 2007-01-23 | 21 | 30 | 27.5 | 18.4 | 2.5 | 2.6 | 5577.5 | 1310.7125 | 265
HS_154R | RP68 | Sector 10 | 2007-01-31 | 19 | 28 | 27.5 | 18.4 | 0.5 | 0.6 | 5539 | 1301.665 | 25E
HS_155L | LP61 | Sector 10 | 2007-01-30 | 19 | 27 | 27.3 | 18.4 | -0.3 | 0.6 | 5519.75 | 1297.1412 | 25B
HS_156L | LP64 | Sector 10 | 2007-01-31 | 19 | 21 | 26 | 18.3 | -5 | 0.7 | 5404.25 | 1269.9987 | 246
HS_157R | RP30 | Sector 10 | 2007-02-01 | 19 | 28.6 | 26.5 | 18.6 | 2.1 | 0.4 | 5550.55 | 1304.3792 | 260
HS_158L | LP68 | Sector 10 | 2007-02-02 | 20.5 | 27.5 | 28 | 18.6 | -0.5 | 1.9 | 5529.375 | 1299.4031 | 25C
HS_159R | RP28 | Sector 10 | 2007-02-12 | 19 | 27 | 25.5 | 18.4 | 1.5 | 0.6 | 5519.75 | 1297.1412 | 25B
HS_160L | LP69 | Sector 10 | 2007-02-12 | 18 | 30 | 28 | 18 | 2 | 0 | 5577.5 | 1310.7125 | 265
HS_161L | LP65 | Sector 1 | 2007-02-20 | 19 | 29 | 28.3 | 18.2 | 0.7 | 0.8 | 5558.25 | 1306.1888 | 262
HS_162L | LP71 | Sector 1 | 2007-02-19 | 19 | 28 | 27.6 | 19 | 0.4 | 0 | 5539 | 1301.665 | 25E
HS_163L | LP72 | Sector 1 | 2007-02-20 | 20 | 30 | 28.8 | 18 | 1.2 | 2 | 5577.5 | 1310.7125 | 265
|  |  |  | T Read SW (MCM) | T Read SW (Pixel Bus) | T Read HW_Ext_Chain (Pixel Bus) | T chuck | DT (Pixel Bus) | DT (MCM) | R EXT_Chain (Calculated) | V ADC (Calculated) | HEX from Rudholf
Iap=0,235 mA
Rextchain=5000+19,25T
T= (Rextchain-5000)/19,25